## Slide 1
Introduction to Light Microscopy
Peter Evennett
peter@microscopical.co.uk

## Slide 2
The Objective Lens
Modern microscope objective lenses are marked with several pieces of important information
The standard marking is
Magnification/Numerical Aperture
In the picture above: 20x/0.60 and 40x/1.0
Find the magnifications and apertures of your microscope lenses
The Aperture is the measure of the fineness of detail the lens is able to pick up – and is what you pay for!
It is the aperture that makes Magnification worthwhile

## Slide 3: The importance of Objective Aperture
Objective
lens
Information
collected
Information
wasted
and some rays
– and some information –
will NOT be collected,
and will be wasted.
Some of these rays
- and some information -
will be collected by the objective
Every ray leaving the object carries some information about fine detail in the object
Resolution depends on angular aperture
of objective

## Slide 4: Microscope Illumination
Once you have bought the objective lenses, there is little you can do to improve your microscope image…
…but it can easily be made worse by poor illumination of the specimen

## Slide 5
The standard method of microscope illumination is Köhler Illumination
devised by
August Köhler
in 1893

## Slide 6: Köhler Illumination enables us to -
Light up the specimen uniformly
  over an adjustable area
    for objectives of different magnifications
    And at the same time to
Illuminate the objective aperture uniformly
  over an adjustable angle
    for objectives of different apertures

## Slide 7
Condenser lens
Stage
Objective lens
Eyepiece
Image focus knobs
Coarse & Fine
Condenser Focus
Field Diaphragm
Condenser Aperture diaphragm
Parts of the microscope
Identify these parts on your microscope
Note: many of the microscopes in the lab do not have a Field Diaphragm

## Slide 8
The diaphragms are NOT intended for adjusting the brightness of the image
The Field Diaphragm sets the area of specimen illuminated, and is adjusted according to magnification
What are the diaphragms for?
The
Condenser Aperture Diaphragm sets the angle of the cone of light illuminating the specimen, and is adjusted according to objective NA

## Slide 9: Setting up Köhler Illumination
Start with both diaphragms (if you have two) wide open
Lamp turned on
Condenser lens raised to top
Lowest magnification objective in place
A stained slide on the stage
  and get a focused image

## Slide 10
Stained slide on microscope, and image brought into focus

## Slide 11
Partly close the Field Diaphragm
the one on the base of the microscope
(not all the microscopes
in the lab have one)
until area illuminated is reduced

## Slide 12
You are now seeing the slide, together with an image of the field diaphragm at the edge, and which is probably
- not sharply focused on specimen
(has fuzzy edge)
- and not centred on field of view
Fuzzy edge

## Slide 13
Adjust condenser focus (up and down) until image of Field Diaphragm is sharply focused on specimen
- but probably still not centred on field of view
Sharp edge
Field of view

## Slide 14
Adjust condenser centring screws until image of Field Diaphragm is centred on field of view.
Then open field diaphragm to illuminate full field

## Slide 15
Take out an eyepiece and look down microscope tube
You are seeing the full aperture of objective
Full aperture of objective

## Slide 16
Full aperture of objective

## Slide 17
75%
100%
Close condenser aperture diaphragm until it is
Iluminating about 75% of the full aperture of objective

## Slide 18
Do not close Illuminating Aperture too far
This will probably increase contrast
But this will be at the expense of resolving the fine detail
75% of
full aperture

## Slide 19: Illuminating Aperture
Correct
Too small

## Slide 20
Illuminating aperture correct

## Slide 21
Illuminating aperture too small

## Slide 22
Illuminating aperture much too small

## Slide 23: Köhler Illumination provides
Control of Area illuminated by the
Illuminated Field Diaphragm,
which is adjusted according to magnification.
Control of Angle of illumination by the
Illuminating Aperture Diaphragm
(the condenser diaphragm),
which is adjusted according to objective aperture.

## Slide 24: A short cut when checking many slides
Set the field diaphragm so that the whole of the field of view is illuminated when using the low-power objective.
  This is so that you can scan the slide and find the area of interest that you’ll want to study at higher power.
Then, having changed to the higher power objective, set the condenser aperture diaphragm correctly – c75% of objective NA
  This is necessary to provide optimum illumination for the larger-aperture high-power objective which you’ll be using for detailed study of the slide. This is the most important factor.
This procedure will result in adequate (though not quite perfect) illuminating conditions for both low and high power images.
  For maximum image quality when necessary (eg for photography), repeat all adjustments to recommended settings.

## Slide 25
Illuminating aperture wide open:
much larger than
objective aperture
Gradually close illuminating aperture diaphragm.
Illuminating aperture now equal to
objective aperture
Removing stray light
Optimum setting:
Stray light removed but
still have adequate
illuminating aperture
Image
brightness
Illuminating aperture
much smaller than
objective aperture
Reducing illuminating aperture
X
A short cut for setting the Illuminating aperture diaphragm
Poor image!

## Slide 26
For a simple demonstration and explanation of microscope optics
Go to https://www.youtube.com/watch?v=60_jgZtyR6U
Or Google
Understanding the light microscope Peter Evennett

## Slide 27
Principles of Light Microscopy
Peter Evennett
peter@microscopical.co.uk